# Appendix A: Measurement Quality Objective Tables

- Table A1: Measurement Quality Objectives* - Conventional Analytes in Fresh and Marine Waters

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 10 analytical runs | 80-120% recovery |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Per 20 samples or per analytical batch, whichever is more frequent | 80-120% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 80-120% recovery |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent (chlorophyll: n/a) | 80-120% recovery RPD<25% for duplicates |
| Laboratory Duplicate | Per 20 samples or per analytical batch, whichever is more frequent (chlorophyll: per method) | RPD<25% (n/a if native concentration of either sample<RL) |
| Internal Standard | Accompanying every analytical run as method appropriate | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | RPD<25% (n/a if native concentration of either sample<RL) |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

- Table A2: Measurement Quality Objectives* – Conventional Analytes in Fresh and Marine Waters – Total, Fixed, Suspended and Volatile Solids

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Laboratory Duplicate | Per 20 samples or per analytical batch, whichever is more frequent. No duplicate required for SSC | RPD<25% (n/a if native concentration of either sample<RL) |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | RPD<25% (n/a if native concentration of either sample<RL) |
| Field Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

- Table A3 (DRAFT): Measurement Quality Objectives* – Bacteria and Pathogens

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration | Check temperatures in incubators twice daily with a minimum of 4 hours between each reading | Per analytical method or manufacturer’s specifications |
| Filter Sterility Check | Perform one filter sterility check each day samples are analyzed; unless otherwise specified by method | No growth on filter |
| Laboratory Blank | Per batch | <RL1 |
| Laboratory Positive Control Samples | Per culture medium or reagent lot | ≥ RL1 |
| Laboratory Negative Control Samples | Per culture medium or reagent lot | < RL1 |
| Laboratory Duplicate | Per 10 samples or per analytical batch, whichever is more frequent unless otherwise specified by method | Within 95% confidence interval (as defined by IDEXX Laboratories) To be updated with text from focus group |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count (coliforms: one per 25 tube dilution tests) | Within 95% confidence interval (as defined by IDEXX Laboratories) To be updated with text from focus group |
| Field Blank, Travel Blank, Equipment Blank | Per method | < RL1 |

*Unless method specifies more stringent requirements

1RL – Reporting Limit

2RPD – Relative Percent Difference

- Table A5: Measurement Quality Objectives* – Inorganic Analytes in Fresh and Marine Waters

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 10 analytical runs | 80-120% recovery |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Per 20 samples or per analytical batch, whichever is more frequent | 75-125% recovery (70-130% for MMHg) |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 75-125% recovery (70-130% for MMHg) |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | 75-125% recovery (70-130% for MMHg); RPD<25% |
| Laboratory Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% (n/a if native concentration of either sample<RL) |
| Internal Standard | Accompanying every analytical run when method appropriate | 60-125% recovery |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | RPD<25% (n/a if native concentration of either sample<RL), unless otherwise specified by method |
| Field Blank, Equipment Blank | Per method | Blanks<RL for target analyte |

*Unless method specifies more stringent requirements

- Table A6: Measurement Quality Objectives* – Volatile Organic Compounds in Fresh and Marine Waters

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 12 hours | RF for SPCCs same as initial calibration; RF of CCVs must be within 20% of initial calibration |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Method Validation: as many as required to assess accuracy and precision of method before routine analysis of samples; Routine Accuracy Assessment: per 20 samples or per analytical batch (preferably blind) | 70-130% recovery if certified; otherwise 50-150% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 50-150% recovery, or based on 3x the standard deviation of laboratory's actual method recoveries |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% |
| Laboratory Duplicate | Per method | Per method |
| Surrogate or Internal Standard | Per method | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | Per method |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

- Table A7: Measurement Quality Objectives* – Semi-Volatile Organic Compounds in Fresh and Marine Waters

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 12 h | RF for SPCCs same as initial calibration; RF of CCVs must be within 20% of initial calibration |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analyte |
| Reference Material | Method Validation: as many as required to assess accuracy and precision of method before routine analysis of samples; Routine Accuracy Assessment: per 20 samples or per analytical batch (preferably blind) | 70-130% recovery if certified; otherwise, 50-150% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 50-150% recovery, or based on 3x the standard deviation of laboratory's actual method recoveries |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% |
| Laboratory Duplicate | Per method | Per method |
| Surrogate or Internal Standard | Per method | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | Per method |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analyte |

*Unless method specifies more stringent requirements

- Table A8: Measurement Quality Objectives* – Synthetic Organic Compounds in Fresh and Marine Waters

| Laboratory Quality Control | Frequency of Analysis | Measurement Quality Objective |
| --- | --- | --- |
| Calibration Standard | Per analytical method or manufacturer’s specifications | Per analytical method or manufacturer’s specifications |
| Continuing Calibration Verification | Per 10 analytical runs | Water: 85-115% recovery Sediment: 85-115% recovery Tissue: 75-125% |
| Laboratory Blank | Per 20 samples or per analytical batch, whichever is more frequent | <RL for target analytes |
| Reference Material | Method Validation: as many as required to assess accuracy and precision of method before routine analysis of samples; Routine Accuracy Assessment: per 20 samples or per analytical batch (preferably blind) | 70-130% recovery if certified; otherwise, 50-150% recovery |
| Matrix Spike | Per 20 samples or per analytical batch, whichever is more frequent | 50-150% recovery, or based on 3x the standard deviation of laboratory's actual method recoveries |
| Matrix Spike Duplicate | Per 20 samples or per analytical batch, whichever is more frequent | RPD<25% |
| Laboratory Duplicate | Per method | Water: RPD<25% (n/a if native concentration of either sample<RL) Sediment: Per method Tissue: Per method |
| Surrogate or Internal Standard | Per method | Per method |
| Field Quality Control | Frequency of Analysis | Measurement Quality Objective |
| Field Duplicate | 5% of total project sample count | Per method |
| Field Blank, Travel Blank, Equipment Blank | Per method | <RL for target analytes |

* Unless method specifies more stringent requirements. ELISA results must be assessed against kit requirements

- Table A25: Measurement Quality Objectives* - Field Measurements in Fresh and Marine Waters**

| Water Quality Parameter | Recommended Device | Units | Resolution | Reporting Limit | “Electronic Specs” Accuracy** |
| --- | --- | --- | --- | --- | --- |
| Depth | Stadia Rod/Staff Gauge | m | 0.01 | 0.02 | n/a |
| Dissolved Oxygen | Polarographic or Luminescence Quenching | mg/L | 0.1 | 0.2 | ± 0.2 |
| pH | Electrode | None | 0.1 | n/a | ± 0.2 |
| Salinity | Refractometer or Conductivity Cell | ‰ | 2 | 2 | ± 2 |
| Specific Conductivity | Conductivity Cell | µS/cm | 1 | 2 | ± 2 |
| Temperature | Thermistor or Bulb | °C | 0.1 or 0.5 | n/a | ± 0.1 |
| Total Chlorophyll | Optical Fluorescence Chlorophyll Probe | µg/L | 0.1 | n/a | n/a |
| Transmissivity | Optical probe | % | 0.005 | 0.01 | +/- 5.0 % |
| Turbidity | Portable Turbidimeter or Optical Probe | NTU | 1 | 5 | ± 1 |
| Velocity | Flow Meter | ft/s | 0.05 | 0.1 | Follow manufacturer’s instructions |

* Unless method specifies more stringent requirements

** This table may not include all field analyses or be applicable Conductivity Temperature Depth Profiler (CTD). Please refer to method or manufacturer instructions for guidance.