Data Acceptability Criteria for Analysis of Water Quality Samples

for Bacteria--Pathogen Indicators

| QA Procedure | Parameter | Frequency1 | Criterion | Corrective Action |  |
| --- | --- | --- | --- | --- | --- |
| Coliform and Enterococcus Bacteria Analyses |  |  |  |  |  |
| Field Blanks | Contamination | 1 per event | < TRL or < sample ÷ 5 |  | Examine field log. Identify contamination source. Qualify data as needed. |
| Lab Method Blanks (Sterility Checks) | Contamination | 1 per batch | < TRL |  | Identify contamination source. Clean equipment and slides. Check reagents. Re-analyze blank. |
| Lab Duplicate | Precision2 | 1 per 10 samples, and at least 1 per batch | Rlog≤ 3.27•mean RLog |  | Recalibrate and reanalyze. |
| Lab Negative Control Samples | Contamination | 1 per culture medium or reagent lot | < TRL |  | Identify source. Clean equipment and prepare new media. Re-examine negative control. |
| Lab Positive Control Samples | Assay function | 1 per culture medium or reagent lot | ≥ TRL |  | Identify and correct problem. Re-examine positive control. |
| Assess percent of data successfully collected | Data Completeness | 1 per planned sample event | 90% |  | Reschedule sample events as necessary or appropriate. |

Notes: TRL = Target Reporting Limit

(1) The method documentation defines an analytical batch as an “uninterrupted series of analyses”.

(2) Rlog is the absolute difference between logarithms of coliform counts for duplicate analyses. The mean Rlog is determined by performing duplicate analyses on the first 15 positive samples analyzed for each matrix type.