Data Acceptability Criteria for analysis of sediment total organic carbon and sediment grain size.

| QA Sample | QA Measure | Minimum Frequency | Criteria | Corrective Action |
| --- | --- | --- | --- | --- |
| Method Blank | Contamination by reagents, laboratory ware, etc. | One per batch. Grain size: N/A | < MDL or < 30% of lowest sample | Identify and eliminate contamination source. Reanalyze all samples in batch. Qualify data as needed. |
| Certified Reference Material | Accuracy | TOC: every 15 samples. Grain size: N/A. | Within 95% confidence interval of the certified value | Review raw data quantification reports. Check instrument response using calibration standard. Recalibrate and reanalyze CRM and samples. Repeat analysis until control limits are met. |
| Replicates | Precision | One per batch | RSD < 20% precision | Check calculations and instruments. Recalibrate and reanalyze. If problem persists, then identify and eliminate source of imprecision and reanalyze. |
| Laboratory control material (LCM) | Accuracy & Precision | One per batch of 20 or fewer samples. Grain size: N/A | Within 20–25% consensus value | Review raw data quantification reports. Check instrument response using calibration standard. Recalibrate and reanalyze CRM and samples. Repeat analysis until control limits are met. |

MDL = method detection limit; RPD = relative percent difference; RSD = relative standard deviation

Sediment TOC

Blanks and a reference material supplied by the instrument manufacturer, Coulometrics, Inc.  will be analyzed a minimum of three times daily during sample analysis.

Grain Size

Standard reference materials will be analyzed with every batch of samples.  These include NIST SRM 1003b glass spheres and a narrow-sized garnet standard supplied by the instrument manufacturer.  In addition, at least one sample in twelve will be analyzed in duplicate to determine precision.  The precision criterion is ±20%.