| HYDROLAB MULTIPROBE CALIBRATION/MAINTENANCE LOG |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Calibration ____ |  | Post Calibration ____ |  | Initials: |
| Date: | Time: |  | Instrument: |  | Battery Voltage: |  |
| If this is a post calibration, give date of original calibration ______ |  |  |  |  |  |  |
| Function |  | Temp. of Standard | Value of Standard | Initial Reading | Calibrated To | Comments |
| Specific Conductance |  |  |  |  |  |  |
| pH calibrated (~7) |  |  |  |  |  |  |
| pH slope (~ 4/10) |  |  |  |  |  |  |
| Dissolved Oxygen |  |  |  |  |  |  |
| DATA NEEDED FOR DISSOLVED OXYGEN CALIBRATION |  |  |  |  |  |  |
| Altitude (A )=______________feet above msl |  |  | Barometric Pressure _________ inches |  |  |  |
| Barometric Pressure (BP) Options |  |  | Barometirc Pressure Formulas |  |  |  |
| Barometer |  |  | Barometric Pressure (inches) ________ x 25.4 = BP ________mm |  |  |  |
| From Local Source After Correction (CBP) |  |  | BP _________ mm = CBP _______mm - 2.5 (altitude ____/100) |  |  |  |
| Estimated From Altitude Only |  |  | BP _________ mm= 760 mm - 2.5 (altitude _____/100) |  |  |  |
| For Older Hydrolabs: Table DO Value______ x ALTCORR______ x BAROCORR ______= DO Standard _______ |  |  |  |  |  |  |
|  |  | Calibration ____ |  | Post Calibration ____ |  | Initials: |
| Date: | Time: |  | Instrument: |  | Battery Voltage: |  |
| If this is a post calibration, give date of original calibration ______ |  |  |  |  |  |  |
| Function |  | Temp. of Standard | Value of Standard | Initial Reading | Calibrated To | Comments |
| Specific Conductance |  |  |  |  |  |  |
| pH calibrated (~7) |  |  |  |  |  |  |
| pH slope (~ 4/10) |  |  |  |  |  |  |
| Dissolved Oxygen |  |  |  |  |  |  |
| DATA NEEDED FOR DISSOLVED OXYGEN POST CALIBRATION |  |  |  |  |  |  |
| Barometric Pressure (BP) Options |  |  | Barometirc Pressure Formulas |  |  |  |
| Barometer |  |  | Barometric Pressure (inches) ________ x 25.4 = BP ________mm |  |  |  |
| From Local Source After Correction (CBP) |  |  | BP _________ mm = CBP _______mm - 2.5 (altitude ____/100) |  |  |  |
| Estimated From Altitude Only |  |  | BP _________ mm= 760 mm - 2.5 (altitude _____/100) |  |  |  |
| For Older Hydrolabs: Table DO Value______ x ALTCORR______ x BAROCORR ______= DO Standard _______ |  |  |  |  |  |  |
| Check previous maintenance and use; do the following before calibration: |  |  |  |  |  |  |
| Polish conductivity electrodes. Must be polished within the last two months or once every 15 field trips |  |  |  | Date: | Name/Comments: |  |
| Change pH reference probe solution. Must be renewed within last two months or once every 15 field trips. |  |  |  | Date: | Name/Comments: |  |
| Inspect D.O. membrane for nicks or bubbles. Must be changed within last six months or once every 15 field trips. |  |  |  | Date: | Name/Comments: |  |
| Change battery in 400 series sonde. Change once a year. Change internal batteries for newer generation products according to guidelines in product manual. |  |  |  | Date: | Name/Comments: |  |

| YSI MULTIPROBE CALIBRATION/MAINTENANCE LOG |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Calibration |  |  |  |  |  |  |
| Date: | Time: |  | Employee Name: |  |  |  |
| Battery Voltage: |  |  | Sonde Type and Serial No. |  |  |  |
| Function |  | Temp. of Standard | Value of Standard | Initial Reading | Calibrated To | Comments |
| Specific Conductance (high) ~100 mS/cm |  |  |  |  |  |  |
| Specific Conductance (med) ~10 mS/cm |  |  |  |  |  |  |
| Specific Conductance (low) ~1 mS/cm |  |  |  |  |  |  |
| Conductivity Cell Constant (see *) |  |  |  |  |  | Range 4.6 to 5.45 |
| pH calibrated ( at pH 7) |  |  |  |  |  |  |
| pH mV for pH 7 solution |  |  |  |  |  | Range 0 to 40 mV |
| pH calibrated (at pH 4 10 – circle one) |  |  |  |  |  |  |
| pH mV for pH 10 pH mV for pH 4 |  |  |  |  |  | Range: -160 to 40 mV Range: +160 to 40 mV |
| Dissolved Oxygen (% sat) |  |  |  |  |  |  |
| Dissolved Oxygen Charge |  |  |  |  |  | Range 25 to 75 |
| Dissolved Oxygen Gain (see *) |  |  |  |  |  | Range 0.8 to 1.7 |
| DATA NEEDED FOR DISSOLVED OXYGEN CALIBRATION |  |  |  |  |  |  |
| Altitude (A )=______________feet above msl |  |  | Barometric Pressure _________ inches |  |  |  |
| Barometric Pressure (BP) Options |  |  | Barometirc Pressure Formulas |  |  |  |
| Barometer |  |  | Barometric Pressure (inches) ________ x 25.4 = BP ________mm |  |  |  |
| From Local Source After Correction (CBP) |  |  | BP _________ mm = CBP _______mm - 2.5 (altitude ____/100) |  |  |  |
| Estimated From Altitude Only |  |  | BP _________ mm= 760 mm - 2.5 (altitude _____/100) |  |  |  |
| Deployment Checklist (required for data logging only) |  |  |  |  |  |  |
| Logging Interval: Yes No |  |  | SDI-12 Autosleep enabled: Yes No |  |  |  |
| DO warm-up time: |  |  | Battery Volts in Sonde (days): |  |  |  |
| RS 232 Autosleep enabled: Yes No |  |  | Available Memory in Sonde (days): |  |  |  |
| Post-Calibration |  |  |  |  |  |  |
| Date: | Time: |  | Employee Name: |  |  |  |
| Battery Voltage: |  |  | Sonde Type and Serial No. |  |  |  |
| Function |  | Temp. of Standard | Value of Standard | Initial Reading | Comments |  |
| Specific Conductance |  |  |  |  |  |  |
| pH calibrated (~7) |  |  |  |  |  |  |
| pH slope (~ 4/10) |  |  |  |  |  |  |
| Dissolved Oxygen |  |  |  |  |  |  |
| Check previous maintenance and use; do the following before calibration: |  |  |  |  |  |  |
| Brush conductivity cell. Must be cleaned within the last two months or once every 15 field trips |  |  |  | Date: | Name/Comments: |  |
| Inspect D.O. membrane for nicks or bubbles. Must be changed within last six months or once every 15 field trips. |  |  |  | Date: | Name/Comments: |  |
| Comments: |  |  |  |  |  |  |

DO Saturation Values

Oxygen Content of Air-Saturated Freshwater at 760 mm Hg

| TempC | Oxygen Solubility at 760 mmHg |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.0 | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 |
| 1 | 14.17 | 14.13 | 14.09 | 14.06 | 14.02 | 13.98 | 13.94 | 13.90 | 13.87 | 13.83 |
| 2 | 13.79 | 13.75 | 13.72 | 13.68 | 13.65 | 13.61 | 13.57 | 13.54 | 13.50 | 13.47 |
| 3 | 13.43 | 13.40 | 13.36 | 13.33 | 13.29 | 13.26 | 13.22 | 13.19 | 13.15 | 13.12 |
| 4 | 13.08 | 13.05 | 13.01 | 12.98 | 12.94 | 12.91 | 12.88 | 12.84 | 12.81 | 12.77 |
| 5 | 12.74 | 12.71 | 12.68 | 12.64 | 12.61 | 12.58 | 12.55 | 12.52 | 12.48 | 12.45 |
| 6 | 12.42 | 12.39 | 12.36 | 12.33 | 12.30 | 12.27 | 12.23 | 12.20 | 12.17 | 12.14 |
| 7 | 12.11 | 12.08 | 12.05 | 12.02 | 11.99 | 11.96 | 11.93 | 11.90 | 11.87 | 11.84 |
| 8 | 11.81 | 11.78 | 11.75 | 11.73 | 11.70 | 11.67 | 11.64 | 11.61 | 11.59 | 11.56 |
| 9 | 11.53 | 11.50 | 11.48 | 11.45 | 11.42 | 11.40 | 11.37 | 11.34 | 11.31 | 11.29 |
| 10 | 11.26 | 11.23 | 11.21 | 11.18 | 11.15 | 11.13 | 11.10 | 11.07 | 11.04 | 11.02 |
| 11 | 10.99 | 10.97 | 10.94 | 10.92 | 10.89 | 10.87 | 10.84 | 10.82 | 10.79 | 10.77 |
| 12 | 10.74 | 10.72 | 10.69 | 10.67 | 10.64 | 10.62 | 10.60 | 10.57 | 10.55 | 10.52 |
| 13 | 10.50 | 10.48 | 10.45 | 10.43 | 10.41 | 10.39 | 10.36 | 10.34 | 10.32 | 10.29 |
| 14 | 10.27 | 10.25 | 10.23 | 10.20 | 10.18 | 10.16 | 10.14 | 10.12 | 10.09 | 10.07 |
| 15 | 10.05 | 10.03 | 10.01 | 9.98 | 9.96 | 9.94 | 9.92 | 9.90 | 9.87 | 9.85 |
| 16 | 9.83 | 9.81 | 9.79 | 9.77 | 9.75 | 9.73 | 9.71 | 9.69 | 9.67 | 9.65 |
| 17 | 9.63 | 9.61 | 9.59 | 9.57 | 9.55 | 9.53 | 9.51 | 9.49 | 9.47 | 9.45 |
| 18 | 9.43 | 9.41 | 9.39 | 9.37 | 9.35 | 9.34 | 9.32 | 9.30 | 92.8 | 9.26 |
| 19 | 9.24 | 9.22 | 9.20 | 9.19 | 9.17 | 9.15 | 9.13 | 9.11 | 9.10 | 9.08 |
| 20 | 9.06 | 9.04 | 9.02 | 9.01 | 8.99 | 8.97 | 8.95 | 8.93 | 8.92 | 8.90 |
| 21 | 8.88 | 8.86 | 8.85 | 8.83 | 8.81 | 8.80 | 8.78 | 8.76 | 8.74 | 8.73 |
| 22 | 8.71 | 8.69 | 8.68 | 8.66 | 8.65 | 8.63 | 8.61 | 8.60 | 8.58 | 8.57 |
| 23 | 8.55 | 8.53 | 8.52 | 8.50 | 8.49 | 8.47 | 8.45 | 8.44 | 8.42 | 8.41 |
| 24 | 8.39 | 8.38 | 8.36 | 8.35 | 8.33 | 8.32 | 8.30 | 8.29 | 8.27 | 8.26 |
| 25 | 8.24 | 8.23 | 8.21 | 8.20 | 8.18 | 8.17 | 8.15 | 8.14 | 8.12 | 8.11 |
| 26 | 8.09 | 8.08 | 8.06 | 8.05 | 8.03 | 8.02 | 8.01 | 7.99 | 7.98 | 7.96 |
| 27 | 7.95 | 7.94 | 7.92 | 7.91 | 7.89 | 7.88 | 7.87 | 7.85 | 7.84 | 7.82 |
| 28 | 7.81 | 7.80 | 7.78 | 7.77 | 7.76 | 7.75 | 7.73 | 7.72 | 7.71 | 7.69 |
| 29 | 7.68 | 7.67 | 7.65 | 7.64 | 7.63 | 7.62 | 7.60 | 7.59 | 7.58 | 7.56 |
| 30 | 7.55 | 7.54 | 7.52 | 7.51 | 7.50 | 7.49 | 7.47 | 7.46 | 7.45 | 7.43 |
| 31 | 7.42 | 7.41 | 7.40 | 7.38 | 7.37 | 7.36 | 7.35 | 7.34 | 7.32 | 7.31 |
| 32 | 7.30 | 7.29 | 7.28 | 7.26 | 7.25 | 7.24 | 7.23 | 7.22 | 7.20 | 7.19 |
| 33 | 7.18 | 7.17 | 7.16 | 7.15 | 7.14 | 7.13 | 7.11 | 7.10 | 7.09 | 7.08 |
| 34 | 7.07 | 7.06 | 7.05 | 7.03 | 7.02 | 7.01 | 7.00 | 6.99 | 6.97 | 6.96 |
| 35 | 6.95 | 6.94 | 6.93 | 6.92 | 6.91 | 6.90 | 6.88 | 6.87 | 6.86 | 6.85 |

Table of pH Calibration Standards

| Temperature (C) | pH 4 Standard | pH 7 Standard | pH 9 Standard | pH 10 Standard |
| --- | --- | --- | --- | --- |
| 10 | 4.00 | 7.06 | 9.33 | 10.15 |
| 11 | 4.00 | 7.06 | 9.32 | 10.14 |
| 12 | 4.00 | 7.05 | 9.31 | 10.13 |
| 13 | 4.00 | 7.05 | 9.30 | 10.12 |
| 14 | 4.00 | 7.04 | 9.29 | 10.11 |
| 15 | 4.00 | 7.04 | 9.28 | 10.10 |
| 16 | 4.00 | 7.04 | 9.27 | 10.09 |
| 17 | 4.00 | 7.03 | 9.26 | 10.08 |
| 18 | 4.00 | 7.03 | 9.25 | 10.07 |
| 19 | 4.00 | 7.02 | 9.24 | 10.06 |
| 20 | 4.00 | 7.02 | 9.23 | 10.05 |
| 21 | 4.00 | 7.01 | 9.22 | 10.05 |
| 22 | 4.00 | 7.01 | 9.21 | 10.04 |
| 23 | 4.00 | 7.00 | 9.20 | 10.03 |
| 24 | 4.00 | 7.00 | 9.19 | 10.01 |
| 25 | 4.00 | 7.00 | 9.18 | 10.00 |
| 26 | 4.00 | 7.00 | 9.17 | 9.99 |
| 27 | 4.00 | 7.00 | 9.16 | 9.98 |
| 28 | 4.01 | 6.99 | 9.16 | 9.98 |
| 29 | 4.01 | 6.99 | 9.15 | 9.97 |
| 30 | 4.01 | 6.99 | 9.14 | 9.96 |
| 31 | 4.01 | 6.99 | 9.13 | 9.95 |
| 32 | 4.01 | 6.99 | 9.13 | 9.94 |
| 33 | 4.02 | 6.98 | 9.12 | 9.93 |
| 34 | 4.02 | 6.98 | 9.12 | 9.92 |
| 35 | 4.02 | 6.98 | 9.11 | 9.91 |
| 36 | 4.02 | 6.98 | 9.10 | 9.90 |
| 37 | 4.02 | 6.98 | 9.09 | 9.89 |
| 38 | 4.02 | 6.98 | 9.08 | 9.88 |